**Transair (Karomed )- changing the mattress onto transport mode**

|  |
| --- |
|  |

Advice Sheet No. 8

In the event of power failure change the mode of the bed using the instructions below for transport mode.

**Transport Mode**

If it is necessary to move the bed with the pump disconnected it is possible to retain the air in the mattress by simply joining together the two end connectors of the mattress hose set. Then turn off the pump.



To return to normal mode, disconnect the two end connectors from each other and reattach them to the pump. Then reconnect the pump to the electricity point and turn on. Check that the mattress is in the preferred mode and amend as needed.

**Power fail**

In the event of power loss the alarm will sound continuously until the “ALARM MUTE” button is pressed. Note: whenever power is interrupted, and subsequently restored, the system will reset into **pressurising mode.**

In transport mode the mattress will remain inflated for up to 15 hours depending on the weight of the patient and how promptly the mattress was put into transport mode.

REFERENENCES

Fiona Webber from Karomed .